DOTIFS Fibre Positioning System

User Guide — Grid Scan & Position Recording

System Design & Development: Vishal Jain

Project Coordinator: Deepa Modi

April 2026

# 1. Accessing the Application

Open a Chrome or Edge browser and navigate to:

dotifs.netlify.app

The application opens with three tabs: GCS Commands (terminal), Live Dashboard, and Grid Scan. Ensure the PI controllers are powered on and connected via USB-Serial before proceeding. The computer to which it is connected should have essential USB-to-serial drivers.

[Screenshot: Application home page with tabs visible]

# 2. Connecting to PI Controllers

Step 1: Select the correct baud rate from the dropdown in the header (default: 9600)

- Set Bits to  of all controllers with the one connected to the computer using USB cable, should not have the address bit of 1

- Also Ensure controllers are connected in a daisy chain.

- Ensure proper power supply to controllers and stages

Step 2: Click "Connect" — a browser dialogue will ask you to select the serial port

Step 3: Choose the COM port connected to the PI controller and click "Connect"

Step 4: The connection indicator turns green when connected

Step 5: Go to the GCS Commands tab and click "Scan IDs for *IDN?" to detect all controllers in the daisy chain

[Screenshot: Connected state with scan results showing controller IDs]

# 3. Grid Scan — Initial Setup

Click on the "Grid Scan" tab (password protected). After entering the password:

Step 1: Set X controller ID and Y controller ID from the dropdowns (e.g., X = ID 1, Y = ID 2/3)

Step 2: Set the Axis number (default: 1)

Step 3: Set File units to "um (micron)" — this ensures coordinates are correctly converted to mm for the stage

Step 4: Check "Invert X" and/or "Invert Y" if the stage moves opposite to what you expect under the microscope

Step 5: Click "Load .txt positions" and select the MLA coordinate file (e.g., MLA_2_I1_3_rot_cropped.txt)

Step 6: The 12x12 hex grid will populate with the loaded coordinates

[Screenshot: Grid Scan tab with loaded coordinates and toolbar settings]

# 4. Setting Your Origin (Grid 0,0 Reference)

Before any grid movement, you must set the origin. The origin defines where grid position (0,0) is on the physical stage. All other grid points will move relative to this position.

Step 1: Click "Set Your Origin" button in the toolbar

Step 2: The origin panel opens showing live X and Y positions

Step 3: Use the step size presets (0.1, 1, 10, 100, 1000 um) or enter a custom value

Step 4: Use the arrow buttons to manually move the stage while looking through the microscope

Step 5: Align the fibre insertion needle precisely over the first hole (grid position 0,0)

Step 6: Click "Lock as Origin" — the current position is saved as the reference origin

Step 7: The status bar shows "Origin: X=... Y=... " with a lock icon

Step 8: Click "Close" to return to the grid

Alternatively: Click "FRF Both > Lock as Origin" to home both axes using the FRF command and automatically set that position as the origin.

[Screenshot: Set Your Origin panel with jog controls and Lock button]

# 5. Moving to Grid Points (Fibre Insertion Mode)

With the origin set, you can now navigate to each fibre position on the grid:

Step 1: Click any dot on the hex grid — a confirmation dialog appears

Step 2: The dialog shows the exact commands that will be sent:
    - MOV to origin (both axes)
    - Wait for on-target
    - MVR by the file offset (relative move from origin)

Step 3: Click "Move" to execute — the stage first goes to origin, then MVR to the target point

Step 4: Wait for "Arrived" status in the status bar

Step 5: Use the Fine Adjust (nudge) arrows on the right panel to make small corrections while looking through the microscope

Step 6: Mark the point: "Yes" (fibre inserted) or "No" (no fibre) or "Delete/Skip"

Step 7: Click "Next Point" to automatically move to the next unvisited point in serpentine order

[Screenshot: Move confirmation dialog showing command preview]

[Screenshot: Grid with fibre/no-fibre/skipped states and nudge controls]

# 6. Fine Adjust (Nudge Controls)

After the stage arrives at a grid point, use the nudge controls on the right panel to fine-tune the position while viewing through the microscope:

Step 1: Enter step size in um (e.g., 1, 5, 65, 75, 250) or use preset buttons

Step 2: Enter number of steps (e.g., step = 65 um, steps = 4, total = 260 um)

Step 3: The "Total" line shows the exact distance that will be moved

Step 4: Click the arrow buttons: Up/Down for Y axis, Left/Right for X axis

Step 5: Each click sends one MVR command — the stage moves by the total amount

Step 6: The last sent command is shown at the bottom for verification

# 7. Recording Positions (Calibration Mode)

Record Mode allows you to capture the actual stage position at each grid point. This creates a calibration file that can be reused for identical MLA assemblies.

## 7.1 How to Record

Step 1: Set your origin first (Section 4)

Step 2: Click "Record: OFF" button in the toolbar — it toggles to "Record: ON" (amber)

Step 3: Using nudge control  move the stage to a fibre hole using the nudge controls

Step 4: Look through the microscope and align precisely

Step 5: Click the corresponding grid cell (the dot on the hex grid)

Step 6: A confirmation dialog appears: "Record position for point #N?"

Step 7: Click "Record" — the software reads POS? from both axes

Step 8: The position is stored relative to your origin (in um)

Step 9: The grid dot turns amber/yellow to show it has been recorded

Step 10: Repeat for each fibre position you want to record

Step 11: The status bar shows "Rec: N" — the count of recorded points

[Screenshot: Record Mode ON with some amber/yellow recorded points on grid]

## 7.2 Exporting Recorded Positions

Step 1: Click "Export Recorded .txt" button in the toolbar

Step 2: A .txt file downloads with only the recorded points

Step 3: The file includes the origin reference position in the header

Step 4: File format: col  row  X_um  Y_um (same as input format)

Step 5: Only points that were actually recorded are included (not all 144)

## 7.3 Loading a Recorded File

Step 1: Click "Load .txt positions" and select a previously recorded file

Step 2: The software detects the origin header and shows a dialog:

Step 3:     "Use Same Origin" — sets origin to the recorded values (same physical setup)

Step 4:     "Set New Origin" — opens the origin panel to set a new reference (different setup)

Step 5: Recorded points appear as amber/yellow dots on the grid

Step 6: You can now click these points to move the stage to those recorded positions

[Screenshot: Origin choice dialog when loading recorded file]

# 8. Quick Reference

## Grid Point Colors

Grey: Unvisited — not yet moved to

Amber (pulsing): Current — stage is at this point

Green (with tick): Fibre inserted successfully

Dark grey: No fibre — visited but skipped

Red (with X): Deleted/skipped — auto-scan will skip this point

Amber (with border): Recorded — position has been captured

## Key Buttons

Set Your Origin — Open panel to manually position and lock the grid origin

Go to Origin — Move both axes back to the locked origin position

FRF Both XY — Home both axes using reference switches

Next Point — Auto-move to the next unvisited point in serpentine order

Record: ON/OFF — Toggle between move mode and position recording mode

Export Recorded .txt — Download recorded positions as a coordinate file

Invert X / Invert Y — Flip axis direction to match microscope view

# Stagger Calculation:

handles this correctly. When generating the 144 points:

if(row % 2 === 0){

// even rows (0,2,4...): aligned with origin — no stagger

x = col * rowPitch;

} else {

// odd rows (1,3,5...): offset by stagger from (0,1) measurement

x = stagger + col * rowPitch;

}

The stagger direction (left or right) is automatically determined from your recorded positions. When you record (0,0) and (0,1):

stagger = (0,1).x - (0,0).x

- If (0,1) is to the left of (0,0) → stagger is negative

- If (0,1) is to the right of (0,0) → stagger is positive

The code uses the actual measured difference, so it captures the real direction. It doesn't assume left or right — it uses exactly what the stage reports.

So if your grid is:

Row 0 (even):    ●  ●  ●     ← origin here

Row 1 (odd):   ●  ●  ●       ← shifted left (negative stagger)

Row 2 (even):    ●  ●  ●

Row 3 (odd):   ●  ●  ●       ← shifted left again

The stagger will come out negative, and all odd rows will correctly shift left. No manual configuration needed — the 3 measurements tell everything.

Prepared by: Vishal Jain

Project Coordinator: Deepa Modi

Project: DOTIFS — Devasthal Optical Telescope Integral Field Spectrograph

Application: dotifs.netlify.app

Date: April 2026