DOTIFS QuickGrid

User Guide

System Design & Development: Vishal Jain

Project Coordinator: Deepa Modi

May 2026

# 1. What QuickGrid Does

QuickGrid is the simplest way to drive fibre placement into the DOTIFS Micro-Lens Array (MLA). It guides the operator through four clear steps, manages stage positioning, and tracks the fibre status for every hole in the 12 x 12 grid.

It is designed for one-handed routine work: open the page, click through the four steps, and start placing fibres. Advanced settings stay hidden until needed.

[Screenshot: QuickGrid main view]

# 2. Opening QuickGrid

Open the DOTIFS application in a browser (Chrome or Edge), then click the "QuickGrid" tab at the top of the page. The page shows four step circles at the top (Connect, Set Origin, Load File, Ready). Each circle turns green as the step is completed.

# 3. The Four Steps

## Step 1 — Connect

Click "Connect". A browser dialog asks which serial port to use:

Pick the COM port that is connected to the PI stage controllers

Click "Connect" in the dialog

The step circle turns green when connected

[Screenshot: Connect step, port selection dialog]

## Step 2 — Set Origin

The Origin is the physical location of hole (0,0) on the MLA — the reference point that all other holes are measured from. Getting the Origin right is the most important step.

Procedure:

Use the nudge arrows (right side panel) to move the stage to hole (0,0)

Look through the microscope and align the fibre insertion needle exactly over the hole

Verify by attempting to physically insert the fibre — the fibre should enter cleanly

Only after the fibre enters successfully, click "Set this as Origin"

A prompt appears: "Do you really want to set this as the grid origin (0,0)?"

Click "Yes, Lock" — the step circle turns green

Important: A small error in the origin (even 5 micrometers) will offset ALL 144 holes by the same amount. Spend extra time on this step. Physically inserting the fibre into hole (0,0) is the most reliable way to confirm alignment.

[Screenshot: Set Origin prompt with X, Y values]

## Step 3 — Load File

Click "Load File" and choose the MLA coordinate file. Two types of files are supported, and QuickGrid detects which one automatically:

### Pixel File (from image processing)

This is the original file containing the centroid pixel coordinates of each hole detected from the MLA image. When you load this:

QuickGrid converts each pixel coordinate to physical micrometers using the formula:

distance (microns) = sqrt(dx^2 + dy^2) x 3.45

The factor 3.45 is the camera scale (microns per pixel)

All 144 hole positions are calculated relative to your Set Origin

A new position file is automatically downloaded with a timestamp, ready for reuse

### Position File (previously saved)

This is a file that was saved earlier from QuickGrid — already in micrometers, no conversion needed. When you load this:

QuickGrid reads the saved Origin from the file header

A prompt appears asking: "Use Same Origin" or "Use Current Origin"

Choose "Use Same Origin" if the MLA is in the exact same physical mount as before

Choose "Use Current Origin" if you have remounted the MLA at a new location

If the file contains recorded corrections from previous work, they are restored automatically

[Screenshot: File load and origin choice dialog]

## Step 4 — Ready

Once all three steps above are green, you can start moving across the grid and inserting fibres. The grid shows all 144 holes as circles in their hex-staggered arrangement matching the physical MLA.

# 4. Working With the Grid

## Click a Hole to Move

Click any circle on the grid to move the stage to that hole. A confirmation dialog shows exactly which commands will be sent and where the stage will end up. Click Move to proceed.

## Fine-Align Under the Microscope

After the stage arrives, use the nudge controls in the right panel to make small corrections until the fibre needle is precisely over the hole:

Enter a step size in micrometers (try 1 for fine, 5 or 10 for coarser)

Use the arrows: up, down, left, right

Each click moves the stage by the chosen step size

Watch the live X, Y position update in the panel

## Mark Fibre Status

After placing (or failing to place) a fibre, mark the hole status in the right panel:

Yes — fibre inserted successfully (hole turns green with a tick)

No — no fibre placed here (hole turns dark grey)

Skip / Delete — exclude this hole from auto-scan (hole turns red with an X)

## Record the Correction

After fine-aligning, click "Record Position" to save the small correction you made. QuickGrid stores how far you nudged from the calculated position. Next time you visit this hole (or load this file later), the stage will go directly to the corrected position without needing to re-align manually.

Recorded holes appear with an amber ring on the grid. The status bar shows how many holes have been recorded.

[Screenshot: Point details panel with Record button and Yes/No/Skip]

# 5. Saving and Reloading Your Work

## Download Position File

Click "Export Recorded" at the top to download a position file containing the calculated positions plus any recorded corrections. The file name includes a timestamp.

Use this file:

To resume work later — load it back and pick up where you left off

To process a second identical MLA — load it and use the same corrections

As a permanent record of where each fibre was placed

## Reload a Position File

Use "Load File" (Step 3) and pick the saved position file. QuickGrid asks whether to use the same origin as before or your currently locked origin. Any recorded corrections in the file are restored.

# 6. Other Useful Buttons

| Button | What it does |
| --- | --- |
| Speed (mm/s) | Sets the universal stage speed for all moves and nudges. Default 0.1 mm/s is safe for microscope work. |
| Apply (speed) | Sends the speed to both X and Y controllers immediately. |
| Acknowledge All Errors | Reads error codes from both stages and clears them. Use this if a stage reports an error. |
| Export Recorded | Downloads the current position file with calculated + recorded values. |
| Reset Grid | Clears all 144 points and their states. Origin stays locked. Confirmation required. |

# 7. Tips for Best Results

Always confirm the Origin by physically inserting a fibre into hole (0,0) before locking it.

Use 0.1 mm/s speed during fine alignment so movements are slow and easy to follow under the microscope.

Record the correction at each hole — it helps with the next MLA and prevents re-doing the alignment work.

Keep the Acknowledge All Errors button in mind. If something gets stuck, it usually clears the issue.

Save your work often by clicking Export Recorded. You can always reload that file to continue.

If a hole has no fibre or is bad, mark it Skip / Delete so auto-scan moves past it.

# 8. Hole Colours on the Grid

| Colour | Meaning |
| --- | --- |
| Grey | Unvisited — not yet moved to |
| Blue ring | Loaded — coordinates are ready, not yet visited |
| Amber (glowing) | Current — stage is at this point right now |
| Green with tick | Fibre placed successfully |
| Red with X | Skipped or deleted by user |

Prepared by: Vishal Jain

Project Coordinator: Deepa Modi

Project: DOTIFS

Date: May 2026